## Sheet1
|  | Revision |  |  | Description of Change
V1
Item Number | Function | Risk Analysis | [6 empty] | Risk Control |  |  |  |  | Risk/Benefit Analysis | Information/ Comments
| Functional Output | HAZARD (Potential cause of Hazard) | Hazardous Situation | HARM (Potential adverse effect) | Cause | SEVERITY | OCCURRENCE | RPN | RISK MITIGATION | SEVERITY | OCCURRENCE | RPN | Risk reduced as far as possible? (afap) | Benefits Outweigh Risks? (Yes/No) | Comments or information be be provided to User
Description |  | Description of why the product will not perform conform specs |  | Description of resulting injury, damage |  |  |  |  | The way by which the risk is reduced/ eliminated |  |  |  | afap | Yes/No | e.g. IFU, labeling
Example | Debris or Fragment | Small fragment from device created during use | Fragment is swallowed and in GI tract | Fragment causes blockage requiring surgery | Material strength is too low and not designed for impact |  |  |  | Material Specification of high impact polystyrene
1 | Debris or Fragment | Small fragment from device created during use | Baby puts small fragment in their mouth | Fragment causes choking and subsequently asphyxiation which could result in death | Improper device assembly, incorrect device assembly instructions, user error | 5 | 3 | 15 | Specification of high impact material and manufacturing of pieces no smaller than 1 in | 5 | 2 | 10 | afap | Yes | Refer to IFU
2 | Device Soothing Design | Device endangers baby's ability to breathe | Baby's mouth or nose covered by device | Suffocation of the baby which could result in death | Improper instructions and user error | 5 | 3 | 15 | Device has a large 14 inch opening for baby's face | 5 | 1 | 5 | afap | Yes | Refer to IFU
|  | Device could be give rise to SIDS | Device causes baby to roll on stomach |  | Baby not properly secured in device or function to secure baby malfunctions | 5 | 3 | 15 | Fail-safe mechanism to prevent baby from rolling over | 5 | 2 | 10 | afap | Yes | Refer to IFU
3 | Device Material | Device material retains checmical irritants during cleaning of device | Release of chemical irritants | Skin Irritation to the baby | Improper instructions and user error on how to clean device which can be retained by material | 2 | 4 | 8 | Use of material the minimizes retaining of cleaning chemicals | 2 | 3 | 6 | afap | Yes | Refer to IFU
|  | Device could be irritable to the skin | Friction on skin |  | Improper instruction and device sizing | 2 | 5 | 10 | Avoidance of synthetic fibers and parts shall individually pass topical irritation test with 100% negative results | 2 | 3 | 6 | afap | Yes | Refer to IFU
4 | Electrical Functions | Device has the potential to electrocute | Electrical shock when plugging in device | Electrocution of the baby and/or parent | Faulty electrical wiring of device and insultation of electrical components | 4 | 2 | 8 | Airtight seal electrical wiring | 4 | 1 | 4 | afap | Yes | Refer to IFU
|  | Device has the potential to start a fire | Small outlet fire | Burns to the baby and/or parent | Faulty outlet connection | 5 | 2 | 10 | Include baby-proof outlet covers with product | 5 | 1 | 5 | afap | Yes | Refer to IFU
FINAL |  | Overall Residual Risk is Acceptable (Yes/No): | [11 empty] | Yes
Severity Scale
| 5 | Catastrophic | Can result in death
| 4 | Critical | Can result in permanent impairment or life-threatening injury
| 3 | Serious | Can result in injury or impairment requiring professional medical intervention
| 2 | Minor | Can result in temporary injury or impairment not requiring professional medical intervention
| 1 | Negligible | Can result in inconvenience or temporary discomfort
Occurence Scale
| 5 | Frequent | ≥1 in 1,000
| 4 | Probable | 1 in 10,000 ≤X<1 in 1,000
| 3 | Occasional | 1 in 100,000 ≤X<1 in 10,000
| 2 | Remote | 1 in 1,000,000 ≤X<1 in 100,000
| 1 | Improbable | < 1 in 1,000,000