## Sheet1
Tracebility Matrix | Project Name:Baby Soothers | DHF: Date 14MAR2022 | Page 1 of 1
Item # | User Needs Intended Uses | Design Inputs (Technical Requirement) | Acceptance Criteria Ideal/Goal Limit |  | Design Outputs | Design Verification (Outputs = Inputs) | Design Validation (User Needs Met)
1 | Customer Requirements (e.g. Functional)
|  |  | Ideal Criteria | Limit Criteria
1.1 | The device is simple to assemble and operate. | Device shall take a person on average 30 min to set up | 30 min | 1 hr
|  | Device shall have no more than 1 button to power on and off the device | 1 | 2
|  | Device requires 0 screws to assemble | 0 | 1
|  | Device shall have no more than 5 removable parts | 5 | 7
1.2 | The device is inexpensive. | Device shall cost no more than $200 retail | 200 | 250
|  | Device shall take no more than $50 to manufacture | 50 | 75
1.3 | The device does not wake parent(s) up when in use. | Device does not exceed 25 dB of sound. | 25 dB | 50 dB
|  | Device brightness level does not exceed 20 lumens | 20 lumens | 30 lumens
1.4 | The device is travel-friendly. | Device shall weigh less than 2 lbs | 2 lb | 2.5 lbs
|  | Device compacts to less than 22 in x 14 in x 9 to fit in a carry-on bag | 10 in x 10 in x 5 in | 15 in x 12 in x 7 in
|  | Device does contains 0% iron, nickel, cobalt, aluminum, copper, or brass so as to not set off metal detectors. | 0% | 0%
|  | Device does not contain more than 3.4 liquid ounces so as to pass through TSA. | 3.4 oz | 3.4 oz
1.5 | The device is washable. | Device shall be no more than 20 in x 14 in x 9 in dimensions when unfolded | 20 in x 13 in x 8 in | 22 in x 14 in x 9 in
|  | Device shall contain at least 60% cotton and 40% polyester | 60% cotton and 40% polyester | 50% cotton and 50% polyester
|  | Device shall be washed with hot water for 30 minutes and dried on low tumble for 30 minutes. | Wash with hot water for 30 minutes and dry on low tumble for 30 minutes | Wash with hot water for 30 minutes and dry on low tumble for 45 minutes
1.6 | The device lasts for a duration of infancy. | Device's battery shall last for approximately 2 years before it's unable to recharge and its function is impaired | 2 years | 1.5 years
1.7 | The device does not interfere with feeding and diaper changes. | Device has a 14 inch opening for baby's face. | 14 in | 12 in
|  | Device has a 20 inch opening at the waist for access to waist and down. | 20 in | 15 in
2 | Performance Requirements (e.g. Mechanical)
2.1 | The device soothes the baby. | Device shall decrease the amount of baby crying at night by 50% | 50% | 25%
2.2 | The device is not loud. | Device will not exceed 25 dB of sound | 25 dB | 50 dB
|  | Device shall send text message alert to parents phone when battery falls below 20% or power input of less than 100 volts is detected. | Phone alert below 20% power and 100 volts | Phone alert below 10% power and 100 volts
2.3 | The device notifies parent(s) if the baby does not stop crying or fussing. | Device shall send text message alert to parents phone after baby has cried for more than 15 minutes | 15 min | 30 min
2.4 | The device is rechargeable. | Device shall contain 1.5 inch opening for removal of battery pack during washing | 1.5 in | 1.3 in
|  | Device contains 2.75 in x 0.75 in x 1.5 in battery pack for 2 AA batteries. | 2.75 in x 0.75 in x 1.50 in | 2.69 in x 0.73 in x 1.31 in
|  | Device takes 2 hours to reach full charge | 2 hours | 3 hours
|  | Device battery lasts 12 hours for a single full charge | 12 | 10
|  | Device comes with 120 in length charging cable that plugs into battery pack's USB port. | 120 in | 72 in
3 | Biocompatability Requirements
3.1 | The device does not irritate the skin. | Device parts shall individually pass topical irritation test with 100% negative results | 100% negative | 95% negative
|  | Device parts shall individually pass guinea pig maximization test with 100% negative results | 100% negative | 95% negative
3.2 | The device must not have easily removed parts. | Device shall not have pieces smaller than 1 in | 1 in | 0.9 in
3.3 | The device does not harm the baby. | Device shall not any parts that have under a 2000g BESS score for sharpness | 2000g | 1900g
|  | There is 0 sq inches of exposed wire to prevent electrocution and fires | 0 in^2 | 0 in^2
|  | The device does not contain a section with an area of 14 in x 14 in that could trap the baby's face causing suffocation | 12 in x 12 in | 14 in x 14 in
|  | Device shall not have any parts with a fillet of less than 0.05 in | 0.05 in | 0.04 in
3.4 | The device is nontoxic. | Device parts shall individually pass cytotoxicity assay with greater than 75% cell viability | 75% viability | 75% viability
|  | Device parts shall individually pass Ames genotoxicity test with 100% negative results | 100% negative | 95% negative
4 | Reuglatory Requirements
4.1 | The device must meet the basic standards of this specification as approved by the FDA. | Device meets FDA requirements and is approved to go on the market | Approved | Approved
4.2 | The device must include a visible warning label that clearly states all the warnings involved with the product. | Device comes with warning label which must be 2 in x 2 in with a font size of 12 | 2 in x 2 in and font size 12 | 1 in x 1 in and font size 11
4.3 | The device must not cause any entrapment of the baby. | Device does not exceed 2 lbs | 2 lbs | 2.5 lbs
5 | Interfaces with Other Systems
5.1 | The device is controlled using Bluetooth. | Device's range of bluetooth for the device is 30 sq feet | 30 ft^2 | 20 ft^2
|  | Bluetooth connection is compatable with apple and andriod phones | compatible | compatible
5.2 | The device is compatible with baby monitoring systems. | Device's connects with baby monitor systems within 30 sq feet | 30 ft^2 | 20 ft^2
6 | Other
6.1 | The device comes in a variety of colors and patterns. | Device shall have 10 different color options | 10 options | 5 options
6.2 | The device is customizeable. | Device shall come in 4 different patterns, 10 different colors, and 3 different sizes. | 4 patterns, 10 color options, 3 sizes | 3 patterns, 5 color options, 3 sizes
6.3 | The device comes in a variety of sizes. | Device shall come in 3 different sizes - S (10 in x 10 in x 5 in), M (15 in x 12 in x 7 in), L (22 in x 14 in x 9 ) | S (10 in x 10 in x 5 in), M (15 in x 12 in x 7 in), L (22 in x 14 in x 9 ) | S (10 in x 10 in x 5 in), M (15 in x 12 in x 7 in), L (22 in x 14 in x 9 )