## Sheet1
Tracebility Matrix | Project Name:Baby Soothers | DHF: Date 14MAR2022 | Page 1 of 1
Item # | User Needs Intended Uses | Design Inputs (Technical Requirement) | Acceptance Criteria Ideal/Goal Limit |  | Design Outputs | Design Verification (Outputs = Inputs) | Design Validation (User Needs Met)
1 | Customer Requirements (e.g. Functional)
|  |  | Ideal Criteria | Limit Criteria
1.1 | The device is simple to assemble and operate. | Device shall take a person on average 30 min to set up | 30 min | 1 hr | Instruction Manual | 5 or less pages of instructions | Timed assembly using focus groups
|  | Device shall have no more than 1 button to power on and off the device | 1 | 2 | CAD Drawing | 1 or less buttons in device drawing | Timed usability test using focus groups to assess ability to find button
|  | Device requires 0 screws to assemble | 0 | 1 |  | 0 screws in device drawing | Timed assembly using focus groups
|  | Device shall have no more than 5 removable parts | 5 | 7 |  | 5 or less parts in device drawing | Timed disassembly using focus groups to assess part detachment
1.2 | The device is inexpensive. | Device shall cost no more than $200 retail | 200 | 250 | Part list | Sum of parts bill is $200 or less | Device bill is comparable to market value for other mobiles
|  | Device shall take no more than $50 to manufacture | 50 | 75 | Manufacturing bill | Sum of manufacturing bill is $50 or less | Device bill is comparable to market value for other mobiles
1.3 | The device does not wake parent(s) up when in use. | Device does not exceed 25 dB of sound. | 25 dB | 50 dB | Audio Specification | Sound analysis reports device does not play sounds over 25 dB | Focus group study to assess how often device wakes an average sleeping person
|  | Device brightness level does not exceed 20 lumens | 20 lumens | 30 lumens | Luminous Specification | Lumen analysis reports device brightness does not exceed 20 dB | Focus group study to assess how often device wakes an average sleeping person
1.4 | The device is travel-friendly. | Device shall weigh less than 2 lbs | 2 lb | 2.5 lbs | Weight Specification | Scale analysis does not exceed 2 lbs | Device can fit inside suitcase and is sufficient weight for airplane travel
|  | Device compacts to less than 22 in x 14 in x 9 in to fit in a carry-on bag | 10 in x 10 in x 5 in | 15 in x 12 in x 7 in | CAD Drawing | Dimensional analysis of compacted device does not exceed 22 in x 14 in x 9 in | Device can fit inside suitcase and is sufficient size for airplane travel
|  | Device does contains 0% iron, nickel, cobalt, aluminum, copper, or brass so as to not set off metal detectors. | 0% | 0% | Detailed Part List | Compositional analysis of device contains 0% iron, nickel, cobalt, aluminum, copper, or brass | Device does not set off metal detectors at airport
|  | Device does not contain more than 3.4 liquid ounces so as to pass through TSA. | 3.4 oz | 3.4 oz | Compositional Specification | Compositional analysis of device contains no more than 3.4 liquid ounces | Device does not have enough liquid to alert travel agencies
1.5 | The device is washable. | Device shall be no more than 20 in x 14 in x 9 in dimensions when unfolded | 20 in x 13 in x 8 in | 22 in x 14 in x 9 in | CAD Drawing | Dimensional analysis of unfolded device does not exceed 20 in x 14 in x 9 in dimensions when unfolded | Device can fit inside standard washing machine
|  | Device shall contain at least 60% cotton and 40% polyester | 60% cotton and 40% polyester | 50% cotton and 50% polyester | Composition Specification | Compositional analysis of device contains at least 60% cotton and 40% polyester | Device can handle a standard wash and dry cycle without ruined fabric
|  | Device shall be washed with hot water for 30 minutes and dried on low tumble for 30 minutes. | Wash with hot water for 30 minutes and dry on low tumble for 30 minutes | Wash with hot water for 30 minutes and dry on low tumble for 45 minutes | Part Specification | Bacterial and UV analysis shows no bacterial growth after device is washed for 30 minutes and dried for 30 minutes | Device can handle a standard wash and dry cycle without ruined fabric
1.6 | The device lasts for a duration of infancy. | Device's battery shall last for approximately 2 years before it's unable to recharge and its function is impaired | 2 years | 1.5 years | Package Specification | Battery analysis shows full functionality after a month of 12 hours of usage per day. | Focus group study to assess battery usage by average family
1.7 | The device does not interfere with feeding and diaper changes. | Device has a 14 inch opening for baby's face. | 14 in | 12 in | CAD Drawing | Dimensional analysis of device contains opening of 14 inches or more for baby's face | Focus group analysis of baby feeding ability while device in use
|  | Device has a 20 inch opening at the waist for access to waist and down. | 20 in | 15 in | CAD Drawing | Dimensional analysis of device contains opening of 20 inches or more for baby's waist | Focus group analysis of baby diaper change ability while device in use
2 | Performance Requirements (e.g. Mechanical)
2.1 | The device soothes the baby. | Device shall decrease the amount of baby crying at night by 50% | 50% | 25% | Noise Tracker | Audio analysis reports time per night of baby crying | Focus group to analyze average time a baby cries on a normal night vs while the device is in use
2.2 | The device is not loud. | Device will not exceed 25 dB of sound | 25 dB | 50 dB | Electrical specification | Audio analysis reports device does not exceed 25 dB | Focus group analysis of device audio capabilities
2.3 | The device notifies parent(s) if the baby does not stop crying or fussing. | Device shall send text message alert to parents phone when battery falls below 20% and/or power input of less than 100 volts is detected. | Phone alert below 20% power and 100 volts | Phone alert below 10% power and 100 volts | Electrical (Arduino) specification | Software analysis reports that text message sent after phone battery falls below 20% and/or when power input is less than 100 volts | Focus group with various types of phones to see how long it takes for the texts to arrive from the device
|  | Device shall send text message alert to parents phone after crying of 110-130 dB is detected for more than 15 minutes | 15 min | 30 min |  | Software analysis reports that text message sent when 110-130 dB is detected for more than 15 minutes
2.4 | The device is rechargeable. | Device shall contain 1.5 inch opening for removal of battery pack during washing | 1.5 in | 1.3 in | CAD Design | Dimensional analysis of device reports 1.5 inch opening for battery pack | Focus group for battery and battery pack insertion and removal in device
|  | Device contains 2.75 in x 0.75 in x 1.5 in battery pack for 2 AA batteries. | 2.75 in x 0.75 in x 1.50 in | 2.69 in x 0.73 in x 1.31 in |  | Dimensional analysis of battery pack shows 2.75 in x 0.75 in x 1.5 opening for 2 AA batteries | Focus group for battery and battery pack insertion and removal in device
|  | Device takes 2 hours to reach full charge | 2 hours | 3 hours | Electrical specification | Battery analysis shows that battery takes 2 hours or less to fully recharge when device power is fully depleted | Timed charging of device from moment device with no charge is pulled in unitl battery is fully charged
|  | Device battery lasts 12 hours for a single full charge | 12 | 10 |  | Battery analysis reports that battery lasts 12 hours when device is in full use | Timed running of device in use from full power until it runs out of power.
|  | Device comes with 120 in length charging cable that plugs into battery pack's USB port. | 120 in | 72 in | Material specification | Dimenionsal analysis reports that device USB charging cable is 120 inches in length | Average measurement anaylisis among 100 different device charging chords.
3 | Biocompatability Requirements
3.1 | The device does not irritate the skin. | Device parts shall individually pass topical irritation test with 100% negative results | 100% negative | 95% negative | Material specification | Irritation analysis reports 100% negative results | Skin analysis is run to test irritation levels
|  | Device parts shall individually pass guinea pig maximization test with 100% negative results | 100% negative | 95% negative |  | Guinea pig maximization test reports 100% negative results | Human skin irritation analysis
3.2 | The device must not have easily removed parts. | Device shall not have pieces smaller than 1 in | 1 in | 0.9 in | CAD design | Dimensional analysis of device does not exceed 1 in on the smallest parts | Timed disassembly using focus groups to assess part detachment
3.3 | The device does not harm the baby. | Device shall not any parts that have under a 2000g BESS score for sharpness | 2000g | 1900g | CAD design | Sharpness analysis does not have any parts with under a 2000g BESS score for sharpness | Cadaver test for skin laceration/irritation
|  | There is 0 sq inches of exposed wire to prevent electrocution and fires | 0 in^2 | 0 in^2 | Electrical specifications | Electrical analysis does contain any exposed wires. | Focus group instructed to search for exposed wires on the device
|  | The device does not contain a section with an area of 14 in x 14 in that could trap the baby's face causing suffocation | 12 in x 12 in | 14 in x 14 in | CAD design and drawings | Dimensional analysis of device does not exceed 12 in x 12 in | Baby placement test in device with outside observation
|  | Device shall not have any parts with a fillet of less than 0.05 in | 0.05 in | 0.04 in | CAD design | Dimensional analysis of device does not have fillets of less than 0.05 in on parts | Cadaver test for skin laceration/irritation
3.4 | The device is nontoxic. | Device parts shall individually pass cytotoxicity assay with greater than 75% cell viability | 75% viability | 75% viability | Material specifications | Cytotoxicity assay analysis returns results of 75% viability or higher | Cadaver test followed by toxicity report
|  | Device parts shall individually pass Ames genotoxicity test with 100% negative results | 100% negative | 95% negative |  | Ames genotoxicity test returns 100% negative results | Cadaver test followed by toxicity report
4 | Reuglatory Requirements
4.1 | The device must meet the basic standards of this specification as approved by the FDA. | Device meets FDA requirements and is approved to go on the market | Approved | Approved | FDA specifications | FDA analysis of device returns an approved result | FDA analysis is run
4.2 | The device must include a visible warning label that clearly states all the warnings involved with the product. | Device comes with warning label which must be 2 in x 2 in with a font size of 12 | 2 in x 2 in and font size 12 | 1 in x 1 in and font size 11 | Label specifications | Dimensional analysis of label does not exceed 2 in x 2 in with a 12 point font size | Focus group instructed to read warning label and report their interpretation
4.3 | The device must not cause any entrapment of the baby. | Device does not exceed 2 lbs | 2 lbs | 2.5 lbs | Material selection and CAD drawing | Weight analysis does not exceed 2 lbs | Baby placement test in device with outside observation
5 | Interfaces with Other Systems
5.1 | The device is controlled using Bluetooth. | Device's range of bluetooth for the device is 30 sq feet | 30 ft^2 | 20 ft^2 | Electronics specificatiion | Bluetooth analysis shows that bluetooth can connect up to 30 ft^2 away | Focus group with various types of phones to see if bluetooth conenction is successful
|  | Bluetooth connection is compatable with apple and andriod phones | compatible | compatible |  | Bluetooth analysis shows that bluetooth can connect successfully to devices | Focus group with various types of phones to see if bluetooth conenction is successful
5.2 | The device is compatible with baby monitoring systems. | Device's connects with baby monitor systems within 30 sq feet | 30 ft^2 | 20 ft^2 |  | Connection analysis shows that device can connect with baby monitor systems up to 30 ft^2 away | Focus group with various types of baby monitor systems to see if conenction is successful
6 | Other
6.1 | The device comes in a variety of colors and patterns. | Device shall have 10 different color options | 10 options | 5 options | Material specification | Color review returns 10 different color options | Focus groups are satisfied with the different color options for the device
6.2 | The device is customizeable. | Device shall come in 4 different patterns, 10 different colors, and 3 different sizes. | 4 patterns, 10 color options, 3 sizes | 3 patterns, 5 color options, 3 sizes | Material and manufacturing specification | Design review returns 4 different patterns, 10 different colors, and 3 different sizes | Focus groups are satisfied with the different sizing options for the device and how they fit their child
6.3 | The device comes in a variety of sizes. | Device shall come in 3 different sizes - S (10 in x 10 in x 5 in), M (15 in x 12 in x 7 in), L (22 in x 14 in x 9 ) | S (10 in x 10 in x 5 in), M (15 in x 12 in x 7 in), L (22 in x 14 in x 9 ) | S (10 in x 10 in x 5 in), M (15 in x 12 in x 7 in), L (22 in x 14 in x 9 ) | Manufacturing specification | Dimensional analysis for size "small" does not exceed 10 in x 10 in x 5 in, for size "medium" does not exceed 15 in x 12 in x 7 in, for size "large" does not exceed 22 in x 14 in x 9 in | Focus group with babies of different ages and sizes trying on the various sizes based on what they are expected to fit into