## Sheet1
Tracebility Matrix | Project Name:Baby Soothers | DHF: Date 11APR2022 | Page 1 of 1
Item # | User Needs Intended Uses | Design Inputs (Technical Requirement) | Acceptance Criteria Ideal/Goal Limit |  | Design Outputs | Design Verification (Outputs = Inputs) | Design Validation (User Needs Met)
1 | Customer Requirements (e.g. Functional)
|  |  | Ideal Criteria | Limit Criteria
1.1 | The device is simple to assemble and operate. | Device shall take a person on average 30 min to set up | 30 min | 1 hr | Instructional Video | Test group watches instructional video and assembles device within 30 minutes | Focus group survey to assess difficulty of assembly based on a scale of easy-hard. Assessment shows at least 90% of surveyees say device was easy to assemble.
|  | Device shall have no more than 1 button to power on and off the device | 1 | 2 | CAD Drawing | 1 or less buttons in device drawing
|  | Device requires 0 screws to assemble | 0 | 1 |  | 0 screws in device drawing
|  | Device shall have no more than 5 removable parts | 5 | 7 |  | 5 or less parts in device drawing
1.2 | The device is inexpensive. | Device shall cost no more than $200 retail | 200 | 250 | Part list | Sum of parts bill is $200 or less | Focus group survey of price reasonability. Assessment shows at least 90% of surveyees find product price reasonable.
|  | Device shall take no more than $50 to manufacture | 50 | 75 | Manufacturing bill | Sum of manufacturing bill is $50 or less | Company survey of part price reasonability in comparison to manufactuerer price. Assesment shows at least 90% of surveyees find part prices reasonable in comparison to manufacturer price.
1.3 | The device does not wake caregiver(s) up when in use. | Device does not exceed 25 dB of sound. | 25 dB | 50 dB | Audio Specification | Sound analysis reports device does not play sounds over 25 dB | Focus group study and follow-up survey to assess how often device wakes an average sleeping person. Assesment shows 90% of surveyees find device does not wake them up while they are asleep.
|  | Device brightness level does not exceed 20 lumens | 20 lumens | 30 lumens | Luminous Specification | Lumen analysis reports device brightness does not exceed 20 dB
1.4 | The device is travel-friendly. | Device shall weigh less than 2 lbs | 2 lb | 2.5 lbs | Weight Specification | Weight analysis does not exceed 2 lbs | Focus group packs carry-on with and without device and assesses relative weight. Follow-up survey shows at least 90% of surveyees see none to minimal difference in carry-on weight.
|  | Device compacts to less than 22 in x 14 in x 9 in to fit in a carry-on bag | 10 in x 10 in x 5 in | 15 in x 12 in x 7 in | CAD Drawing | Dimensional analysis of compacted device does not exceed 22 in x 14 in x 9 in | Focus group assesses ease of device fitting inside carry-on regulated bag. Follow-up survey shows 90% of surveyees rate fitting the device inside a carry-on easy on a scale from easy-hard.
|  | Device does contains 0% iron, nickel, cobalt, aluminum, copper, or brass so as to not set off metal detectors. | 0% | 0% | Detailed Part List | Compositional analysis of device shows 0% iron, nickel, cobalt, aluminum, copper, or brass | TSA assessment shows device does not set off metal detectors
|  | Device does not contain more than 3.4 liquid ounces so as to pass through TSA. | 3.4 oz | 3.4 oz | Compositional Specification | Compositional analysis of device contains no more than 3.4 liquid ounces | TSA assessment shows device does not have enough liquid to impede travel through airport
1.5 | The device is washable. | Device attachments and structure supports shall be no more than 20 in x 14 in x 9 in dimensions when unfolded | 20 in x 13 in x 8 in | 22 in x 14 in x 9 in | CAD Drawing | Dimensional analysis of unfolded device does not exceed 20 in x 14 in x 9 in dimensions when unfolded | Focus group assess whether the parts of device can fit inside standard dishwasher. Follow-up survey shows 100% of surveyees were able to fit the device attachments and supports into a standard dishwasher.
|  | Device attachments at least 60% cotton and 40% polyester | 60% cotton and 40% polyester | 50% cotton and 50% polyester | Composition Specification | Compositional analysis shows device contains at least 60% cotton and 40% polyester | Focus groups assess the device attachment durability after 5 standard wash and dry cycles. Follow-up survey shows at least 90% surveyees find zero to minimal fabric degradation.
|  | Device attachments shall be washed with hot water for 30 minutes and dried on low tumble for 30 minutes. | Device attachments are clean after washing 30 minutes and dry on low tumble for 30 minutes | Device attachments are clean after washing with hot water for 30 minutes and dry on low tumble for 45 minutes | Part Specification | Bacterial and UV analysis shows no bacterial growth after device is washed for 30 minutes and dried for 30 minutes | Focus groups assess whether device is clean after a wash cycle with hot water for 30 minutes and dry cycle on low tumble for 30 minutes. Follow-up survey shows at least 90% of surveyees find device clean after wash and dry cycle.
|  | Device structural supports shall be washed with hot water for 30 minutes and air-dried for 30 minutes. | Device structural supports are clean after washing with hot water for 20 minutes | Device structural supports are clean after being washed with hot water for 30 minutes | Part Specification | Bacterial and UV analysis shows no bacterial growth after device is washed | Focus groups assess device structural support cleanliness after being washed with hot water for 30 minutes and air-dried for 30 minutes. Follow-up survey shows at least 90% of surveyees find device structural supports clean after specified washing and drying.
1.6 | The device lasts for a duration of infancy. | Device's battery shall last for approximately 2 years before it's unable to recharge and its function is impaired | 2 years | 1.5 years | Package Specification | Battery analysis shows full functionality after a month of 12 hours of usage per day. | Focus group assess average device battery usage over one week. This avergage shall be scaled up to infacy duration (2 years). Scaled assessment shows at least 90% battery efficiency at 2 years.
1.7 | The device does not interfere with feeding and diaper changes. | Device has a 14 inch opening for baby's face. | 14 in | 12 in | CAD Drawing | Dimensional analysis of device contains opening of 14 inches or more for baby's face | Focus group assess ease of baby feeding while using device. Follow-up survey shows at least 90% of surveyees find that the device does not interfere with baby feeding.
|  | Device has a 20 inch opening at the waist for access to waist and down. | 20 in | 15 in | CAD Drawing | Dimensional analysis of device contains opening of 20 inches or more for baby's waist | Focus group assess ease of a diaper change ability while using device. Follow-up survery shows at least 90% of surveyees find that the device does not interfere with diaper changes.
2 | Performance Requirements (e.g. Mechanical)
2.1 | The device soothes the baby. | Device shall decrease the amount of baby crying at night by 50% | 50% | 25% | Noise Tracker | Audio analysis reports time per night of baby crying | Focus group assesses average number of times a baby cries for one day without using the device vs. for one day using the device. Assessment shows that for at least 90% of participants, baby crying decreased up to 50%.
2.2 | The device notifies caregiver(s) if the baby does not stop crying or fussing. | Device shall send text message alert to caregiver phone when battery falls below 20% and/or power input of less than 100 volts is detected. | Phone alert below 20% power and 100 volts | Phone alert below 10% power and 100 volts | Electrical (Arduino) specification | Software analysis reports that text message sent after phone battery falls below 20% and/or when power input is less than 100 volts | Focus group with various types of phones assess how long it takes for the texts to arrive from the device. Assessment shows that at least 90% of texts take on average 1 minute or less to arrive.
|  | Device shall send text message alert to caregiver phone after crying of 110-130 dB is detected for more than 15 minutes | 15 min | 30 min |  | Software analysis reports that text message sent when 110-130 dB is detected for more than 15 minutes
2.3 | The device is rechargeable. | Device shall contain 1.5 inch opening for removal of battery pack during washing | 1.5 in | 1.3 in | CAD Design | Dimensional analysis of device shows 1.5 inch opening for battery pack | Focus group assess ease of removing battery pack from device. Follow-up survey shows at least 90% of surveyees find removing battery pack from the device easy on a scale of easy-hard.
|  | Device contains 2.75 in x 0.75 in x 1.5 in battery pack for 2 AA batteries. | 2.75 in x 0.75 in x 1.50 in | 2.69 in x 0.73 in x 1.31 in |  | Dimensional analysis of battery pack shows 2.75 in x 0.75 in x 1.5 opening for 2 AA batteries | Focus group assess ease of battery insertion and removal into the battery pack. Assessment shows 100% of participants are able to insert and remove the batteries into the battery pack.
|  | Device takes 2 hours to reach full charge | 2 hours | 3 hours | Electrical specification | Battery analysis shows that battery takes 2 hours or less to fully recharge when device power is fully depleted | Focus group assess device charging when battery is fully depleted over one week. Assessment shows that 100% of devices fully recharged in on average 2 hours or less.
|  | Device battery lasts 12 hours for a single full charge | 12 | 10 |  | Battery analysis reports that battery lasts 12 hours when device is in full use | Focus group assess device battery life for one week. Assessment shows 100% of devices last on average for 12 hours or more for a single full charge.
|  | Device comes with 120 in length charging cable that plugs into battery pack's USB port. | 120 in | 72 in | Material specification | Dimenionsal analysis reports that device USB charging cable is 120 inches in length | Focus group assess ease of charging device with 120 inch length charging cable. Follow-up survey asking whether a 120 inch cable is too short, good length, or too long shows that at least 90% of surveyees believe a 120 inch cable is a good length.
3 | Biocompatability Requirements
3.1 | The device does not irritate the skin. | Device parts shall individually pass topical irritation test with 100% negative results | 100% negative | 95% negative | Material specification | Irritation analysis reports 100% negative results | Negative cadaver test for skin irritation.
|  | Device parts shall individually pass guinea pig maximization test with 100% negative results | 100% negative | 95% negative |  | Guinea pig maximization test reports 100% negative results
3.2 | The device must not have easily removed parts. | Device shall not have pieces smaller than 1 in | 1 in | 0.9 in | CAD design | Dimensional analysis of device does not exceed 1in on the smallest parts | Focus group assess ease of removing parts from device without specific instruction. Assessment shows 100% of participants are unable to remove device parts without instruction.
3.3 | The device does not harm the baby. | Device shall not any parts that have under a 2000g BESS score for sharpness | 2000g | 1900g | CAD design | Sharpness analysis does not have any parts with under a 2000g BESS score for sharpness | Negative cadaver test for skin irritation and laceration.
|  | Device shall not have any parts with a fillet of less than 0.05 in | 0.05 in | 0.04 in | CAD design | Dimensional analysis of device does not have fillets of less than 0.05 in on parts | Negative cadaver test for skin irritation and laceration.
|  | Device contains 0 sq inches of exposed wire | 0 in^2 | 0 in^2 | Electrical specifications | Electrical analysis does not show any exposed wires. | Negative cadaver test for electrocution by device.
|  | Device does not contain any part with an area of 6 in x 6 in or greater that could trap the baby's face or body. | 6 in x 6 in | 7 in x 7 in | CAD design and drawings | Dimensional analysis of device does not exceed 6 in x 6 in | Placement of baby-sized doll in device does not show covering of nose or mouth.
3.4 | The device is nontoxic. | Device parts shall individually pass cytotoxicity assay with greater than 75% cell viability | 75% viability | 75% viability | Material specifications | Cytotoxicity assay analysis returns results of 75% viability or higher | Negative cadaver test followed by toxicity report.
|  | Device parts shall individually pass Ames genotoxicity test with 100% negative results | 100% negative | 95% negative |  | Ames genotoxicity test returns 100% negative results
4 | Regulatory Requirements
4.1 | The device must meet the basic standards of this specification as approved by the FDA. | Device meets FDA requirements and is approved to go on the market | Approved | Approved | FDA specifications | FDA analysis of device returns an approved result | FDA approves the device for sale
4.2 | The device must include a visible warning label that clearly states all the warnings involved with the product. | Device comes with warning label which must be 2 in x 2 in with a font size of 12 | 2 in x 2 in and font size 12 | 1 in x 1 in and font size 11 | Label specifications | Dimensional analysis of label does not exceed 2 in x 2 in with a 12 point font size | Focus group instructed to read warning label and report their interpretation. Follow-up survey shows 100% of surveyees' interpretation of warning label matched warning message.
5 | Interfaces with Other Systems
5.1 | The device is controlled using Bluetooth. | Device's range of bluetooth for the device is 30 sq feet | 30 ft^2 | 20 ft^2 | Electronics specificatiion | Bluetooth analysis shows that bluetooth can connect up to 30 ft^2 away | Focus group with various types of Apple and Android phones assess whether device can be connected to phones using bluetooth. Assessment shows 100% of participants were able to connect their phones to the device within a 30 sq feet range.
|  | Bluetooth connection is compatable with apple and andriod phones | compatible | compatible |  | Bluetooth analysis shows that bluetooth can connect successfully to devices
5.2 | The device is compatible with baby monitoring systems. | Device's connects with baby monitor systems within 30 sq feet | 30 ft^2 | 20 ft^2 |  | Connection analysis shows that device can connect with baby monitor systems within 30 sq feet | Focus groups with various types of compatible baby monitor systems assess whether device can be connected to baby monitor systems via bluetooth. Assessment shows 100% of participants are able to successfully connect their baby monitor systems to the device.
6 | Other
6.1 | The device is customizeable. | Device shall have 10 different color options | 10 options | 5 options | Material specification | Color review returns 10 different color options | Focus groups assess satisfaction with different color options for the device. Follow-up survey shows at least 90% of surveyees are satisfied with color options.
|  | Device shall come in 3 different sizes - S (10 in x 10 in x 5 in), M (15 in x 12 in x 7 in), L (22 in x 14 in x 9 ) | S (10 in x 10 in x 5 in), M (15 in x 12 in x 7 in), L (22 in x 14 in x 9 ) | S (10 in x 10 in x 5 in), M (15 in x 12 in x 7 in), L (22 in x 14 in x 9 ) | Manufacturing specification | Dimensional analysis for size "small" does not exceed 10 in x 10 in x 5 in, for size "medium" does not exceed 15 in x 12 in x 7 in, for size "large" does not exceed 22 in x 14 in x 9 in | Focus groups assess satisfaction with different sizing options for device. Follow-up survey shows at least 90% of surveyees are satisfied with sizing options.