## Matrix
Item # | User Needs Intended Uses | Design Inputs (Technical Requirement) | Acceptance Criteria Ideal/Goal Limit |  | Design Outputs | Essential Design Outputs (Check if essential) | Design Verification (Outputs = Inputs) | Design Validation (User Needs Met)
1 | Customer Requirements (e.g. Functional)
1.1 | Detects motor seizure activity | Detects pressure from hand clenching at a threshold of X N | Detects pressure from hand clenching at a threshold of X N | Detects pressure from hand clenching at a threshold of (X+20) N | Circuit diagram, Arduino IDE code |  | When device is tested hand clenching results in detecting pressure at X N | Users report that seizure is detected when hand is clenched within the threshold
1.2.1 | Alerts user through alarm | Produces beeping alert sound at 80 decibels | Sound produced is 80 decibels | Sound produced is 60 decibels | Assembled glove, Arduino IDE code |  | When device is tested alarm sound is X decibels | Test users report sound is loud enough to hear from outside the room the device is used in
1.2.2 | Alarm alerts nearby caretakers to user condition | Alarm continues to sound after initial threshold met, even after falling below threshold before alarm stopped | Sound produced continuously once threshold is reached | Sound produced for only duration of seizure event, stops without caretaker input at end of event | Assembled glove, Arduino IDE code |  | When device is tested reaching threshold causes sound to go off indefinitely unless physically acknowledged | Test users report that sound does not stop after starting unless physically acknowledged
|  | Alarm shuts off when physically acknowledged and returns alarm to 0 decibels | Sound produced is 0 decibels when button is pressed twice | Sound produced is 0 when button is pressed once | Assembled glove, Arduino IDE code |  | When device is tested pressing button twice in succession causes sound to be reduced to 0 decibels | Test users report that sound stops being produced from device when the intended "off" button is pressed twice in a row
1.3 | Quickly detects when a seizure is happening | Ability to detect seizure within 5 seconds | Seizure detected in 1 second | Seizure detected in 20 seconds | Assembled glove, Arduino IDE code |  | When device is tested meeting seizure criteria results in producing sound in X seconds | Users report that the device alarm sounded within X seconds of meeting the detection criteria/clenching hand
2 | Product Performance Requirements (e.g.Mechanical)
2.1 | Light-weight | Weighs no more than 50 grams | Device weighs 50 grams | Device weighs 80 grams | Assembled glove |  | Final glove and attached parts weigh X grams | Users report that the glove doesn't feel too heavy when worn
2.2 | Device can be worn for long periods of time (e.g. while sleeping) | Stays attached to hand for at least 7 hours | Device stays attached to hand for 8 hours | Device stays attached to hand for 7 hours | Assembled glove |  | Device worn by user stays attached for at least 7 hours | Users report that the device stayed attached while they were using it at night while sleeping
2.3 | Device functions for long periods of time | Has a battery life of at least 7 hours | Device powered by battery for 8 hours | Device powered by battery for 7 hours | Assembled glove |  | Device can function for 7 hours after attaching battery | Users report LED still on in morning on Arduino after use overnight
2.4 | Device can be repetitively used and washed | Electronics can be removed from glove in less than 2 minutes | Electronics can be removed from the glove in 30 seconds | Electronics can be removed from the glove in 2 minutes | Glove, detached parts (pressure sensors, battery, Arduino, buzzer) |  | Users can remove attached parts (pressure sensors, buzzer, breadboard, Arduino) from glove in X seconds | Users report they can easily and quickly remove the device parts attached to the fabric glove
|  | Glove can be removed from hand in 5 seconds | Glove can be removed in 2 seconds | Glove can be removed in 5 seconds | Glove |  | User can remove glove from hand in X seconds | Users report they can quickly and easily remove the fabric glove from their hand
3 | Biocompatability Requirements
3.1 | Safe for use with human skin contact/body, no adverse reaction | No adverse reaction on skin for a minimum of 7 hours when device is in use | Absence of adverse skin reaction for 8 hours | Absence of adverse skin reaction for 7 hours | Material Specification (Link when known, SDS) |  | Conduct trials with users wearing the device for 7 hours and evaluating skin irritation using the "Visual Assessment of Human Skin Irritation" 10 point scale after the time duration | Meets threshold of "0" on "Visual Assessment of Human Skin Irritation" 10 point scale, meaning no pink color on skin after 7 hours
4 | Regulatory Requirements
4.1 | Labeling of directions, intended use, and safety | 90% of users should understand directions, intended use, and safety after reading labels | 100% of users should understand directions, intended use, and safety after reading labels | 90% of users should understand directions, intended use, and safety after reading labels | Labels with instructions for use, safety (Link) |  | Conduct trials of asking users to read and interpret the directions and labeling information | Users report understanding the directions and labeling, are able to explain directions clearly
4.2 | Accuracy of detection validation | Detects 90% of seizures | 100% of seizures detected | 90% of seizures detected | Assembled glove |  | Conduct trials with users wearing device outside of trial for 10 seizure events and determine accuracy | Users report that X out of 10 total seizures were detected by the device
5 | Interfaces with Other Systems
5.1 | Ability to send a separate alert to nearby caretakers' phones | Sends notification of seizure occuring 100% of the time | Notification sent to phone 100% of the time when seizure is occuring | Notification sent to phone 80% of the time when seizure is occuring | Assembled glove, Arduino IDE code |  | Conduct tests with users triggering clenching detection and record alerts received on caretaker phone | Caretakers are able to view phone notifications for seizures on personal phones
5.2 | Ability to track information about the seizure | Records hand clenching for the duration of the seizure in second/minutes | Seizure duration recorded for 100% of seizure events | Seizure duration recorded for 80% of seizure events | Assembled glove, Arduino IDE code |  | Conduct tests with users triggering clenching detection and analyze percentage of seizure events/duration tracked | Users are able to view seizure duration after each seizure event
6 | Other
6.1 | Affordable | Total cost less than 350 dollars with no required subscription plan | Total cost is 100 dollars | Total cost is 150 dollars | Assembled glove, material order forms |  | Cost analysis and manual review of total costs | Review finds that the total material cost for the device is less than 350 dollars
6.2.1 | Able to be used by small-handed people | Fits glove size S-small with vertical hand length 6-6.5 inches | Able to be worn by user with hand length of 6 inches | Able to be worn by user with hand length of 6.5 inches | Assembled glove |  | When device tested on a user with hand length of 6-6.5 inches, assess fit and comfort of glove | Users report that the glove fits and is comfortable
6.2.2 | Able to be used by medium-handed people | Fits glove size M-medium with vertical hand length 6.5-7.5 inches | Able to be worn by user with hand length of 6.5 inches | Able to be worn by user with hand length of 7.5 inches | Assembled glove |  | When device tested on a user with hand length of 6.5-7.5 inches, assess fit and comfort of glove | Users report that the glove fits and is comfortable
6.2.3 | Able to be used by large-handed people | Fits glove size L-large with vertical hand length 7.5-8.5 inches | Able to be worn by user with hand length of 7.5 inches | Able to be worn by user with hand length of 8.5 inches | Assembled glove |  | When device tested on a user with hand length of 7.5-8.5 inches, assess fit and comfort of glove | Users report that the glove fits and is comfortable

## Definitions
Essential Rquirements: Provide essential safety and performance criteria.