## Matrix
Item # | User Needs Intended Uses | Design Inputs (Technical Requirement) | Acceptance Criteria Ideal/Goal Limit |  | Design Outputs | Essential Design Outputs (Check if essential) | Design Verification (Outputs = Inputs) | Design Validation (User Needs Met)
1 | Customer Requirements (e.g. Functional)
1.1.1 | Detects motor seizure activity | Detects pressure from hand clenching from normal force at a threshold of 2.77 psi | Detects pressure from hand clenching at a threshold of 2.77 psi | Detects pressure from hand clenching at a threshold of 0.41 psi | Circuit design, Arduino IDE code |  | Benchtop tests result in hand clenching being detected when a pressure of 2.77 psi is applied to the pressure sensor | Clinical trial of 100 participants report that device detects seizure when hand is clenched within the threshold
1.1.2 |  | Detects pressure from hand clenching from bending finger at a threshold of 200 Ohms (resistance equivalent to full bend of finger) | Detects pressure from finger bending at a threshold of 200 Ohms | Detects pressure from finger bending at a threshold of 300 Ohms |  |  | Benchtop tests result in hand clenching being detected when finger bending of 200 Ohms is applied to the flex sensor sensor
1.2.1 | Alerts user through alarm | Produces beeping alert sound at 80 decibels | Sound produced is 80 decibels | Sound produced is 60 decibels | Circuit design, Arduino IDE code |  | Benchtop tests resulted in an alarm sound at 80 decibels when hand clenching was detected | Clinical trial of 100 participants report that alarm sound is loud enough to hear from outside the room the device is used in
1.2.2 | Alarm alerts nearby caretakers to user condition | Alarm continues to sound after initial threshold met, even after falling below threshold before alarm stopped | Sound produced continuously once threshold is reached | Sound produced for only duration of seizure event, stops without caretaker input at end of event | Circuit design, Arduino IDE code |  | Benchtop tests resulted in alarm sounding indefinitely when threshold was reached, unless physically acknowledged | Clinical trial of 100 participants report that sound does not stop after starting unless physically acknowledged
|  | Alarm shuts off when physically acknowledged and returns alarm to 0 decibels | Sound produced is 0 decibels when button is pressed twice | Sound produced is 0 when button is pressed once | Circuit design, Arduino IDE code |  | Benchtop tests resulted in device alarm reducing to 0 decibels when the button is pressed twice in succession | Clinical trial of 100 participants report that sound stops being produced from device when the intended "off" button is pressed twice in a row
1.3 | Quickly detects when a seizure is happening | Ability to detect seizure within 5 seconds | Seizure detected in 1 second | Seizure detected in 20 seconds | Circuit design, Arduino IDE code |  | Benchtop tests resulted in device alarm producing sound within 3 seconds of seizure sensor threhold/criteria being reached | Clinical trial of 100 participants report that the device alarm sounded within 3 seconds of meeting the detection criteria/clenching hand
2 | Product Performance Requirements (e.g.Mechanical)
2.1 | Light-weight | Weighs no more than 50 grams | Device weighs 50 grams | Device weighs 80 grams | Assembled glove |  | Benchtop weighing of glove and attached parts weigh 50 grams | User survey of 100 participants report that the glove doesn't feel too heavy when worn and are satisfied with the weight of the glove
2.2 | Device can be worn for long periods of time (e.g. while sleeping) | Stays attached to hand for at least 7 hours | Device stays attached to hand for 8 hours | Device stays attached to hand for 7 hours | Assembled glove |  | Benchtop test using silicone casted hand shows that the device stays attached for at least 7 hours with hand movement | Clinical trial of 100 participants report that the device stayed attached while they were using it while sleeping for an average of 6-8 hours overnight
2.3 | Device functions for long periods of time | Has a battery life of at least 7 hours | Device powered by battery for 8 hours | Device powered by battery for 7 hours | Circuit design, assembled glove |  | Benchtop test shows that device can function for 7 hours after attaching battery | Clinical trial of 100 participants report Arduino LED indicating function remaining on after use for 7-8 hours overnight
2.4 | Device can be repetitively used and washed | Electronics can be removed from glove in less than 2 minutes | Electronics can be removed from the glove in 30 seconds | Electronics can be removed from the glove in 2 minutes | Glove, detached parts (pressure sensors, battery, Arduino, buzzer) |  | Benchtop test resulted in attached device parts (pressure sensors, buzzer, breadboard, Arduino) being removed from glove in 30 seconds | User survey of 100 participants report an average 4.5 out of 5 satisfaction score with the time needed and ease of removal of device parts attached to the fabric glove
|  | Glove can be removed from hand in 5 seconds | Glove can be removed in 3 seconds | Glove can be removed in 10 seconds | Glove |  | Benchtop test resulted in removal of glove from hand in 3 seconds | User survey of 100 participants report an average 4.5 out of 5 satisfaction score with the time needed and ease of removal of the fabric glove
3 | Biocompatability Requirements
3.1 | Safe for use with human skin contact/body, no adverse reaction | No adverse reaction on skin for a minimum of 7 hours when device is in use | Absence of adverse skin reaction for 8 hours | Absence of adverse skin reaction for 7 hours | Material Specification (skin-safe T-shirt fabric) |  | Benchtop test resulting in no irritation detected after human epidermis model (mimicking human skin) was in contact/use with the device for 7 hours | Clinical trial with 100 participants wearing the device for 7 hours and reporting an average of "0" for skin irritation using the "Visual Assessment of Human Skin Irritation" 10 point scale after the time duration
4 | Regulatory Requirements
4.1 | Labeling of directions, intended use, and safety | 90% of users should understand directions, intended use, and safety after reading labels | 100% of users should understand directions, intended use, and safety after reading labels | 90% of users should understand directions, intended use, and safety after reading labels | Labels on glove with instructions for use and safety, instruction manual |  | Conduct trials of asking users to read and interpret the directions and labeling information | User test with 100 participants report understanding the directions and labeling and are able to explain directions clearly after reading the labeling
4.2 | Accuracy of detection validation | Detects 90% of seizures | 100% of seizures detected | 90% of seizures detected | Less than 10% false positives when 2.6 psi is applied to device at a frequency of 0.8 second pulses for 5 seconds |  | Benchtop tests with pressure ranging from 0-2.6 psi at a frequency of 0.8 second pulses for 5 seconds resulted in less than 10% false positives | User test of 100 participants using device outside of trial for 10 seizure events report that 9 out of 10 total seizures were detected by the device
5 | Interfaces with Other Systems
5.1 | Ability to send a separate alert to nearby caretakers' phones | Sends notification of seizure occuring 100% of the time | Notification sent to phone 100% of the time when seizure is occuring | Notification sent to phone 80% of the time when seizure is occuring | Arduino IDE code, phone notification |  | Benchtop tests result in a phone notification being sent to caretaker's phone after a seizure is detected | User test of 100 caretakers reported recieving phone notifications for seizures on personal phones for 10 simulated seizures
5.2 | Ability to track seizure duration | Records hand clenching for the duration of the seizure in second/minutes | Seizure duration recorded for 100% of seizure events | Seizure duration recorded for 80% of seizure events | Arduino IDE code, Arduino IoT Cloud Remote app |  | Benchtop tests result in storage of seizure duration for 10 trial seizures after a seizure is detected applying 2.77 psi of pressure | User test of 100 caretakers reported an average 4.5 out of 5 satisfaction score for ease of viewing of seizure duration after each seizure event
6 | Other
6.1 | Affordable | Total cost less than 350 dollars with no required subscription plan | Total cost is 100 dollars | Total cost is 150 dollars | Bill of materials with total cost of $73.77 |  | Cost analysis and manual review of total costs showed that the total cost is less than $100 | User survey of 100 participants reported an average 4.5 out of 5 satisfaction score for ranking the affordability of the device
6.2.1 | Able to be used by users with glove size small | Fits glove size S-small with vertical hand length 6-6.5 inches | Able to be worn by user with hand length of 6 inches | Able to be worn by user with hand length of 6.5 inches | Small glove design |  | Benchtop tests result in glove fitting smoothly on silicon casted hand model with hand length of 6-6.5 inches | User survey of 100 participants who wear glove size small reported an average 4.5 out of 5 satistfaction score for the fit and comfort of the glove
6.2.2 | Able to be used by users with glove size medium | Fits glove size M-medium with vertical hand length 6.5-7.5 inches | Able to be worn by user with hand length of 6.5 inches | Able to be worn by user with hand length of 7.5 inches | Medium glove design |  | Benchtop tests result in glove fitting smoothly on silicon casted hand model with hand length of 6.5-7.5 inches | User survey of 100 participants who wear glove size medium reported an average 4.5 out of 5 satistfaction score for the fit and comfort of the glove
6.2.3 | Able to be used by users with glove size large | Fits glove size L-large with vertical hand length 7.5-8.5 inches | Able to be worn by user with hand length of 7.5 inches | Able to be worn by user with hand length of 8.5 inches | Large glove design |  | Benchtop tests result in glove fitting smoothly on silicon casted hand model with hand length of 7.5-8.5 inches | User survey of 100 participants who wear glove size large reported an average 4.5 out of 5 satistfaction score for the fit and comfort of the glove

## Definitions
Essential Rquirements: Provide essential safety and performance criteria.